# Randolph Community College

## Mammography Equipment Solicitation

### Attachment Packet (A–E)

Solicitation Title: Refurbished Mammography Equipment
Solicitation Number: 864252603
Issuing Agency: Randolph Community College
Division: College Services / Purchasing

# Table of Contents

Attachment A – Refurbishment Certification Form
Attachment B – Installer & Technician Credential Form
Attachment C – Service Provider Disclosure Form
Attachment D – System Equivalency Evaluation Matrix
Attachment E – Manufacturing Date Verification Form

# ATTACHMENT A – REFURBISHMENT CERTIFICATION FORM

Vendor Name: __________________________________________
Equipment Make/Model Offered: __________________________________________
Serial Number(s): __________________________________________

### A. Refurbishment Steps Completed

(Vendor shall check each completed item)

[ ] Full electrical inspection and reconditioning
[ ] Replacement of worn electrical components
[ ] Replacement or refurbishment of detector components
[ ] X-ray generator calibration
[ ] Detector calibration
[ ] Mechanical component restoration
[ ] Cosmetic restoration (covers, panels, shielding surfaces)
[ ] System-level quality control testing
[ ] OEM-equivalent tomosynthesis and CAD performance verification

### B. Required Documentation Attached

(Vendor must attach all items)

[ ] Refurbishment checklist
[ ] Component replacement list
[ ] Electrical safety testing report
[ ] Final QC and calibration results
[ ] Detector performance documentation (uniformity, bad pixel map, flat-field)

### Certification

I hereby certify that the equipment offered meets or exceeds the refurbishment requirements of this solicitation.

Authorized Representative Name: ___________________________________
Title: ___________________________________
Signature: ___________________________________
Date: ___________________________________

ATTACHMENT B – INSTALLER & TECHNICIAN CREDENTIAL FORM

Vendor Name: __________________________________________

### A. Installer Personnel Information

Paste the table below into Word and then Convert Text to Table (Tabs).

| Installer Name | Training/Certification | Type of Certificate | Issuing Entity | Expiration Date | Years of Mammography Installation Experience |
| --- | --- | --- | --- | --- | --- |

(Add rows as needed)

### B. Documentation Attached

[ ] OEM training credentials
[ ] Equivalent manufacturer-recognized training
[ ] Electrician licenses (if applicable)
[ ] Proof of mammography installation experience

### Certification

I certify that the installers listed meet all qualifications required in Sections 3.9 and 3.12 of the solicitation.

Signature: ___________________________________
Date: ___________________________________

# ATTACHMENT C – SERVICE PROVIDER DISCLOSURE FORM

Vendor Name: __________________________________________

### A. Primary Warranty Service Provider

Company Name: __________________________________________
Location: __________________________________________
Distance from RCC: __________ miles
Average On-Site Response Time: __________ hours

### B. Technician Credentials Attached

[ ] OEM or factory training certificates
[ ] Service technician resumes or experience summaries
[ ] Documentation of specialization in mammography systems

### C. Subcontractor Disclosure

If subcontractors will be used for any installation, repair, or maintenance activities, complete the table below.

| Subcontractor Name | Role or Service Provided | Certification Attached (Y/N) | Response Time (Hours) |
| --- | --- | --- | --- |

(Add rows as needed)

### Certification

I certify that the above information is accurate and complete.

Signature: ___________________________________
Date: ___________________________________

# ATTACHMENT D – SYSTEM EQUIVALENCY EVALUATION MATRIX

(To be completed only when offering an “or equal” system)

| Specification Category | IFB Requirement | Vendor Response | Meets Requirement? (Y/N) | Notes |
| --- | --- | --- | --- | --- |
| System Type | Hologic Selenia Dimensions 3D equivalent, 2014 or newer |  |  |  |
| Tomosynthesis | Required |  |  |  |
| 2D FFDM Imaging | Required |  |  |  |
| C-View | Required |  |  |  |
| CAD | Required (ImageChecker + Citra) |  |  |  |
| Detector Type | Selenium TFT |  |  |  |
| Detector Size | 24 × 29 cm |  |  |  |
| Pixel Size | 0.070 mm |  |  |  |
| Spatial Resolution | 7.1 lp/mm |  |  |  |
| Generator | Constant potential HF |  |  |  |
| Tube | Tungsten bi-angular |  |  |  |
| Filters | Rhodium/Silver/Aluminum |  |  |  |
| Workstation Specs | CPU, RAM, GPU, OS |  |  |  |
| DICOM Services | MWL, MPPS, RDSR |  |  |  |
| PACS Compatibility | IQ-View |  |  |  |
| Required Accessories | All listed in specifications |  |  |  |
| Displays | 17” tech monitor; 21.3” 2MP clinical display |  |  |  |
| UPS | Minimum 5-minute backup |  |  |  |
| Refurbishment Documentation | All required documentation included |  |  |  |
| Installation Requirements | Installer qualifications met per Section 3.9 |  |  |  |

### Certification

I certify that this matrix accurately reflects the system offered.

Signature: ___________________________________
Date: ___________________________________

# ATTACHMENT E – MANUFACTURING DATE VERIFICATION FORM

Vendor Name: __________________________________________
Equipment Offered: __________________________________________

### A. Proof of Original Manufacturing Date

Vendor must attach one or more of the following:

[ ] OEM-issued manufacture date statement
[ ] FDA 2579 form
[ ] OEM configuration/build report
[ ] System log showing original build date

### B. Mixed-Component Disclosure

Does the system contain components from multiple manufacturing years?

[ ] NO
[ ] YES — If yes, list each component and its manufacturing year below:

| Component Name | Year Manufactured | Replaced During Refurbishment? (Y/N) |
| --- | --- | --- |

(Add rows as needed)

### C. Certification

I certify that the equipment offered meets the solicitation requirement of being 2014 or newer.

Signature: ___________________________________
Date: ___________________________________